HOW  TO CALIBRATE THE ML11X

This tutorial will show you how calibrate an ML11X on the range

- You will need:

- Calibration range

- ML11X

- Set the instrument on the range with the display facing you.

- Hold the ↑ and turn on the instrument.

- Display will show ‘R0.000X’, release ↑.

- Press MAN button until ‘ERROR’ shows in the display.

- Holding the MAN button use ↑ to step through the display until TYPE shows.

- Release the MAN button.

- Change TYPE from ML11 to ML14

- Turn off unit

- Turn on instrument

- Raise/lower instrument until it is level with the centre of the collimator.

- Using ↑ and ↓ rotate head until beam shines through the collimator

- Hold MAN button and use ↑ to step the display until Y-AXIS shows

- Release MAN button

- Use ↑ and ↓ to change setting until beam is level (on the collimator x-hairs).

- Rotate laser 180˚ and check that beam is level.

- Press MAN button

- Rotate instrument 90˚ and repeat 11-16 for x-axis

- Turn instrument off

- Hold the ↑ and turn on the instrument.

- Display will show ‘R0.000X’, release ↑.

- Press MAN button until ‘ERROR’ shows in the display.

- Holding the MAN button use ↑ to step through the display until TYPE shows.

- Release the MAN button.

- Change TYPE from ML14 to ML11

- Turn off instrument